## Slide 1: Differential pressure measurement device
Basic principles

## Slide 2
Regulated to test pressure
Here is a diagram of the device in “filling” mode
The Reference volume (Ref) is a “zero leak” volume with roughly the same volume as the device under test.
  A welded up flange or pressure system
The Test volume is the device under test, you need to know its volume to the nearest ml
  Easily done by filling with water and taking the weight change. (or use maths)
V1 – V4 are all switching valves, automation is easier
P abs is a normal pressure meter.
Dp is the sensitive differential pressure meter
  Here dp is zero as both volumes are at the same pressure

## Slide 3
Closing V2 stops the supply of test gas
Leave system at least 15 mins to equalise
  Temp changes due to gas decompression/compression.
  Good to also record the Ref volume and Test volumes actual temperatures.
As V1 is open the ref and test volumes act as one so dp is zero.
Can be turned off

## Slide 4
To start testing V2 is closed
Now the Ref. volume and Test volume are separate
Dp will start to change
  Might be an initial wobble due to the volume change done by switching the valve.

## Slide 5
As test gas leaks out of the test volume its pressure decreases.
It’s now different to the reference
As the Ref. has a “zero” leak the change in pressure is only from the leak in the Test volume.
As we know how long V1 has been closed we can get a rate of change of pressure for the test volume mbar/s
We can use the gas laws to determine a mass loss mg/s

## Slide 6
Equalise by opening V1
Tests end when either a time limit is reached or the dp sensor is getting to the maximum, say 500mbar of differential pressure

## Slide 7
Drops due to loss of media

## Slide 8
O

## Slide 9: Why differential pressure sensor?
Normal absolute pressure sensors are not accurate enough
  A meter which can do 60 bar might be good for 0.1 bar increments
  Dp meters can do 0.1 mbar with line pressures 200+ bar
  Measure a small leak rate sooner so the test is quicker.
Use of ref volume removes external influences
  Both volumes are effected equally(ish) to external influences (temp, pressure)

## Slide 10: The maths!
Test volume = 150ml
Gasket length = 200 mm, 0.2m
Test temperature = 20°C
Test media = He (4003 mg/mol)
Pressure lost = 200 mbar
Test time = 20 minutes
Use the gas laws, pV=nRT
  n = pV/RT
Find number of moles,n
p = pressure in Pa = 200 mbar = 20000 Pa
V = volume in m3 = 150ml = 1.5E-4 m^3
R = Ideal gas constant = 8.314 J/mol/K
T = Temp in Kelvin = 20°C = 293.15K
n = 0.001231 mol.
Convert moles to mg
mg He = 0.001231 * 4003 = 4.93 mg He
20min = 1200s so Rate = 4.93/1200 = 0.0041 mg/s
Normalise to length, m 0.0041/0.2 = 0.021 mg/m/s

## Slide 11: Example result – High leak rate
Test is done 3 times to ensure a decent rate can be found
Rate here is 3.6e+1 mg/m/s
  ~100000 ppm (EPA)
190 mbar difference is very quickly hit (10-20 seconds)

## Slide 12: Example result – Medium leak rate
Test is done 3 times to ensure a decent rate can be found
Rate here is 1.0e-1 mg/m/s
  ~500-300 ppm (EPA)
Detected a pressure change of 190mbar in ~35 minutes
  Steady state reached in about 10 mins

## Slide 13: Example result – Low leak rate
Test is done just once
  24 hrs long! (Ta Luft)
Rate here is 1.2e-5 mg/m/s
  Less than 1 PPM (EPA)
Detected next to no pressure change, the extreme limit of this technique
  You can see the dp goes up and down, most likely due to temperature variations
But 10-2 – 10-3 levels are easily possible
100-1 PPM (EPA)

## Slide 14: Example dp meter devices
Keller PD33X – Link 2
Omega PXM80
CS 400
Honeywell A-5
Or just ask your favourite sensor provider
  Requirements – Line pressure greater than 60 bar, range 500 mbar dp.
Note for methane use, you’ll be wanting to get a firesafe version.